The M-124 Ozonometers Network

Arcady Shalamjansky

Main Geophysical Observatory (MGO), St. Petersburg, Russia, October 2002

Instruments:

The Russian network for observing total column ozone consists of 28 stations within Russia and 2(3) Antarctic stations, which are equipped with filter ozonometers, model M-124.  These same instruments were used at stations in Bulgaria (3 stations), Cuba (2), Georgia (1), Latvia (1), Lithuania (1), Kazakhstan (5 ), Tajikistan (1), Turkmenistan (2), Ukraine (4) and Vietnam (3).  This report and schedule refers only to those stations, which are calibrated with assistance from the MGO with some degree of regularity.

The M-124 ozonometer uses two filters to measure the total ozone amount.  The filter spectral intervals have their maximum at 302 and 326 nm each with a half-width of about 20 nm.  With time, the filters could be affected by solarization and this could be the cause of observed systematic errors.

Total ozone is retrieved from direct sun measurements at zenith angles of 20-70 and from blue or cloud zenith measurements at zenith angles of 20- 85 which allows for ozone measurements at northern latitude stations, under practically all weather conditions.

All of the M-124 ozonometers are calibrated against the Dobson spectrophotometer No.108, which is regularly compared with the WMO standard.  Intercomparisons in Boulder (1988), Hradec Králové (1993), Kalavrita (1997) and Hohenpaissenberg (2001) showed that the measurement-scale drift of the Dobson No.108 did not exceed 1.0 %.

# Data quality control methodology

The possibility of the filter characteristic degradation determines the special features of the station equipment and the measurement procedure.  In particular, it is necessary to carry out the regular control of a measurement scale.Each Russian ozone station is equipped with three ozonometers: one dedicated to continuous operation, one spare and a reserve that is sent to the MGO for calibration.

The measurement procedure consists of the mandatory blue or cloud zenith measurements and the direct sun measurements (when sun is clear and the zenith angle in less than 20 70??).  Tables of observation results are sent from the stations to the MGO every month

The MGO control of measurement quality is based on:

- the calibration of M-124 ozonometers against the standard every two years;

- the monthly comparison of two ozonometers, with which any station is equipped.

- the comparison of direct sun and zenith measurements (the difference can not exceed 1-3 %);

- the routine control over the measurements for revealing the "-dependence" which indicates change of instrument spectral characteristics

The MGO analysis is being performed to minimise the systematical errors, which are accounts for about 90 % of the total error.  At the Russian network stations, the error of a single direct sun and zenith ozone measurement by the M-124 ozonometer is no more than 5 %.  After the MGO control and correction procedure, the measurement error could be reduced to 3 %.

Real-time data.

The MGO also receives real-time data each day by telegram and/or e-mail.  These data are then plotted on a chart to compare with satellite ozone measurements (NASA TOMS instrument) and with 300 mb geopotential charts  (as well as 200 mb geopotential and 30 mb temperature charts).  There comparisons are intended to determine main features of total ozone distribution and to find rough (technical) errors in the real-time data.

Telegram real-time data can be used for the WOUDC real-time ozone maps, but are not suitable for the data submission to the WOUDC.

Ozone data for a publication

The MGO strongly recommends that the WOUDC users use only the MGO quality controlled data.  As a minimum requirement, the data should at least be accompanied by comments (qualified or not qualified). It is standard practice for the MGO to submit only qualified data to the WOUDC.

At present, the MGO presents the qualified data only from Russia and Kazakhstan instruments.

Since 1973, the MGO have performed the control of the filter ozonometers (models M-83 and M-124) data for the former USSR ozone network.  The MGO also recommends that the WOUDC users do not use any M-83 data prior to 1973 because of large systematic errors.

State of the M-124 ozonometers network (as of 2002)

The M-124 ozonometers have been in operation for more than 17 years.  Each ozonometer requires regular maintenance and repair.  With assistance from the WMO, a number of M-124 instruments have been upgraded.  Such upgrades include the installation of an inner power supply block, block of thermostabilization and a complete set of new optical filters.

Filter updating is a challenge for the MGO now because Russian industry does not manufacture inexpensive optical glasses anymore.

Up to the year 2002, the MGO upgraded two thirds of the Russian ozonometers.

The instruments located in Lithuania (1 station) and Vietnam (4) were upgraded and calibrated at the MGO in 2002. Also in 2002, two ozonemeters were prepared for Cuba.

The Bulgarian ozonometers were calibrated in 2000.

The Ukraine (3) and Kazakhstan (4) ozonometers were partially upgraded and then calibrated in 1999.

The Latvian ozonometer was calibrated in 1997.

The Georgia, Turkmenistan and Tajikistan instruments have not been calibrated since 1995-96.