## DATA
FLT | DATE | FLIGHT | TYPE | MIN | BOX | MAX | MIN | THERM | I | I | FAULT
No |  |  |  | AIR | TEMP C | HT | P |  | micro A | micro A
[6 empty] | (km) | mB | ISTOR | SYSTEM | DVM
| 1995-11-02 | VIZ1222 | PTU | [7 empty] | OK
| 1995-11-03 | VIZxxxx | PTU | [7 empty] | OK
1 | 1995-11-06 | 1238/951106 | O3 | -66 | 6 | 25 | 6 | YES | -0.06 | 0.03 | OK temp erratic prior to launch
5A0173
2 | 1995-11-08 | 951108 | O3 |  |  | 2 | 774 | YES |  | 0.07 | Fail. Lost signal during ascent, no winds,pressure then temp
|  | Z00801 | [8 empty] | possible interference from Vaisala. Temps erratic at launch
| 1995-11-09 | 951109 | PTU | [7 empty] | No winds, otherwise OK
3 | 1995-11-10 | 951110 | O3 | -64 | -12.8 | 33 | 7 | YES | 0.07 | 0.02 | At 25.5mB jumped to 13.2mB then 7.2 then 45.5
|  | Z00802 | [8 empty] | ozone trace on ascent and descent. No winds, got trapped in
wind co-ordinate plot. Good enough to send to NILU
4 | 1995-11-13 | 951113 | O3 | -64 | 1.7 | 18 | 83 | YES | 0.063 | 0.03 | Lost temp, box temp, RH - RAW data still coming
|  | 5A017305 | [8 empty] | signal strength OK. Bead temp too cold
Scatter in data @T+43 mins. Journal - error computing
sensor resistance, humidity sample rejected, temp
sample rejected
5 | 1995-11-14 | 951114 | O3 | -26 | -2 | 6 | 482 | YES |  | 0.07 | Fail, Bead temp wrong, big jumps, frequency jumped
|  | Z00803 | [8 empty] | then lost temp, bead temp and RH. Journal file, error
computing sensor resistance, temp sample rejected
T+33. RAW temp, RH bead count all 1's. T+10 temp
6 | 1995-11-15 | 951115 | O3 | -75 | -76 | 32 | 7.2 | YES | 0.08 | 0.01 | OK.
Z00804
7 | 1995-11-16 | 951116 | O3 | -26 | 3 | 4 | 603 | YES |  | 0.07 | T+45. Lost temp, bead temp, RH. RAW data coming in but no
|  | Z00805 | [8 empty] | processing. REF counts erratic as on previous occasions
Bead at 7 C at T+7minutes.
11 | 1995-11-17 | 95117 | O3 | -76 | -13.8 | 32 | 6 | YES | 0.139 | 0.2 | OK except we lost several minutes of data when frequency
|  | Z00807 | [8 empty] | rose 1MHz shortly after launch. AFC off. It said match etc but
lost sound of modulation. Returned and data came in again.
Bead temp varying rapidly up and down. -10 C but ozone still
OK. Solution should be frozen and therefore no ozone.
12 | 1995-11-20 | 951120 | O3 | -74 | 12 then | 31 | 8.6 |  | 0.09 | 0.075 | Ht starts at 122M. OK Bead thermistor replaced by resistor. Temp
|  | z00806 |  |  | -6 | [5 empty] | unstable after burst
13 | 1995-11-23 | 951123 | O3 | -74 | 6 | 33.7 | 6.7 |  | 0.06 | 0.026 | OK. Resistor substitute. . box temp 11.9 - 12.8 - 9.0 unstable above
|  | Z00808 | [8 empty] | 10mB. descent 5 C @ 440mb varying by several degrees
14 | 1995-11-27 | 951127 | O3 | -78 | 12 | 27.3 | 16.5 |  |  | 0.025 | OK. 1st attempt at run up failed. No ozone detected outdoors.
|  | Z00810 | [8 empty] | Vizsonde replaced. resistor substitute. Ht starts at 106M
15 | 1995-11-30 | 951130 | O3 | -76 | 12.1 | 35.2 | 4 |  | 0.1 | 0.02 | OK. resistor substitute. signal strength dropped to 100 with AFC
|  | Z00811 | [8 empty] | on after 60 mins. retuned and AFC on, strength returned to almost
max
16 | 1995-12-04 | 951204 | O3 | -75 | 17 | 29 | 11.8 |  | 0.103 | 0.029 | OK. Faulty ozonesonde -plug connection -resistor substitute.
|  | Z00812 | [8 empty] | 2nd ozonesonde OK
17 | 1995-12-06 | 951206 | O3 | -59 | 17 | 19 | 66.1 |  | 0.13 | 0.04 | Fail. Resistor substitute. Signal poor. Flight finished early in mid
|  | Z00813 | [8 empty] | stratosphere, too low for total ozone
18 | 1995-12-08 | 951208 | O3 | -66 | 17 | 33 | 6.7 |  | 0.08 | -0.02 | OK. resistor substitute
Z00814
19 | 1995-12-11 | 951211 | O3 | -56 | 26 to | 29 | 15.3 | YES | 0.11 | 0.053 | Lots of missing data at intervals. Temp and thermistor lost
|  | Z00815 |  |  | -12 | [5 empty] | at about 200mB. Appeared again at about 500mB on descent.
Reprocessed. Total ozone recovered
20 | 1995-12-13 | 951213 | O3 | -67 | 17 | 32.7 | 6.8 |  | 0.08 | 0.028 | Missing data between 190 and 25mB. Temp, Rh Box temp replaced by //
|  | Z00816 | [8 empty] | Resistor substitute. RAW data OK. Reprocessed. Total ozone recovered
Error computing sensor resistance
21 | 1995-12-15 | 951215 | O3 | -68 | 6 | 14.9 | 195 |  | 0.07 | 0.031 | Resistor. First flight in W9000 handbook configuration. Failed
|  | Z00817 | [8 empty] | 195mB box temp sudden drop near trop then fail
Error computing sensor resistance
22 | 1995-12-18 | 951218 | O3 | -81 | 11.9 | 31.8 | 6.7 |  | 0.41 | 0.069 | Resistor. OK. occasional pressure and O3 processed data
|  | Z00818 | [8 empty] | replaced by /////, more so after 40mB.OK
23 | 1995-12-21 | 951221 | O3 | -76.9 | 12 | 22.1 | 33.5 | YES | 0.12 | 0.05 | OK. Thermistor. Balloon burst early, too early for total ozone (213.5DU)
Z00819
24 | 1995-12-30 | 951230 | O3 | -77 | 12.4 | 28.4 | 11.5 |  | 0.05 | 0.104 | Lost signal thrice. Resistor.OK
Z00820
25 | 1996-01-03 | 960103 | O3 | -82 | 13.6 | 30.2 | 9.1 |  | 0.12 | 0.063 | 16K resistor. Twice balloon broken away 30kts. Twice signal lost behind
|  | Z00821 | [8 empty] | shed.OK
26 | 1996-01-05 | 960105 | O3 | -8 | 13.6 | 3.4 | 694 |  | 0.07 | 0.032 | Resistor. Failed, red light on at rear, signal lost 3 times. Couldn't use
|  | Z00822 | [8 empty] | ALT T until T+50 mins. No further data.
27 | 1996-01-08 | 960108 | O3 | -75 | 13.8 | 33.5 | 5.4 |  | 0.07 | 0.018 | Resistor, AFC pulls frequency and signal strength down.OK
Z00823
28 | 1996-01-10 | 960110 | O3 | -80 | 13.9 | 32.3 | 6.3 |  | 0.06 | 0.2 | Resistor. Only 2 LORAN signals.OK
Z00824
29 | 1996-01-12 | 960112 | O3 | -79 | 13.5 | 33.3 | 5.6 |  | 0.07 | 0.1 | Pump didn't start, bad connections, erratic O3 pre-flight, OK after
Z00825
30 | 1996-01-15 | 960115 | O3 | -75 | 14 | 37 | 3.2 |  | 0.09 | 0.046 | Resistor. Initially would not locate, possible battery problem.OK
Z00826
31 | 1996-01-17 | 960117 | O3 | -67 | 3.1 | 23.4 | 35 |  | 0.06 | 0.065 | Resistor. No winds operator error.OK
Z00827
32 | 1996-01-19 | 960119 | O3 | -79 | 13.6 | 32.2 | 7 |  | 0.07 | 0.025 | Resistor. Frequent missing pressure and ozone counts in processed
|  | Z00828 | [8 empty] | data. Up to 4 or 5 per screen.OK
33 | 1996-01-22 | 960122 | O3 | -85 | 16.6 | 32.4 | 6 |  | 0.05 | 0.025 | 14K resistor. OK
Z00829
34 | 1996-01-24 | 960124 | O3 | -79 | 13.6 | 32.1 | 7 |  | 0.05 | 0.015 | Needed reboot to recognize Ser No or LORAN. 16K resistor. OK
Z00830
35 | 1996-01-26 | 960126 | O3 | - | - | - | - | YES |  | 0.059 | FAILED. Lost signal T+570, returned then again at T+1080secs
Z00831
36 | 1996-01-27 | 960127 | O3 | -78 | 14.5 | 33.7 | 5 |  | 0.027 | 0.027 | OK. Resistor. No initial signal, broken batt wire.Freq 406.7 then jumped to
|  | Z00832 | [8 empty] | 408.3. This during run up. MATCH FLIGHT
37 | 1996-01-29 | 960129 | O3 | -82 | 16.6 | 30.4 | 9 |  | 0.06 | 0.01 | OK. Resistor. MATCH FLIGHT
Z00833
38 | 1996-01-30 | 960130 | O3 | -75 | 17.4 | 26.4 | 8 |  | 0.07 | 0.027 | OK. Resistor. Ozone went to zero at 26.5kM so file truncated at that
|  | Z00834 | [8 empty] | height. Pump failure?
39 | 1996-02-02 | 960202 | O3 | -71 | 16.6 | 34.9 | 5.2 |  | 0.1 | 0.036 | Vaisala blotted out Viz signal for 30 secs prior to launch. Ozone profile
|  | Z00835 | [8 empty] | looks odd over 33kM so 14E file truncated. Descent data erratic and lost
@ 217mB. Time of launch not detected by our software
40 | 1996-02-05 | 960205 | O3 | -64 | 17 | 31.1 | 9 |  | 0.77 | 0.053 | OK. Obnormally high system background current
Z00836
41 | 1996-02-06 | 960206 | O3 | -55 | 21 to | 32.6 | 7 | YES | 0.056 | 0.018 | Thermistor. First sonde bead temp, temp and RH all /////. Sonde ser No
|  | Z00837 |  |  | -1.5 | [5 empty] | not transmitted. 2nd sonde OK. Interference from Viz on deck to Vaisala
in flight
42 | 1996-02-08 | 960208 | O3 | -68 | 21 to | 31.5 | 8 | YES | 0.07 | 0.024 | Thermistor. Error 835 METREAD.EXE. Thermistor probably broken.
|  | Z00838 |  |  | 0.2 | [5 empty] | This message appeared ONCE ONLY and temp read OK, maybe some
temps outside of range. No descent data. Good flight
43 | 1996-02-09 | 960209 | O3 | -77 | 14.8 | 31.7 | 7 | YES | 0.12 | 0.014 | Thermistor. Rapid trail off of ozone to zero at burst. Our software
|  | Z00839 | [8 empty] | detected launch at wrong time. Flight truncated at 27194 metres on 14E
file to ensure more accurate total ozone. New total 273.7 Old value 255.7
44 | 1996-02-11 | 960211 | O3 | -78 | 14.8 | 31.8 | 7 |  | 0.09 | 0.026 | Burst detected by our software at 27.2kM, also launch time wrong
|  | Z00840 | [8 empty] | because of 0.7mB drop of pressure on the ground.
45 | 1996-02-12 | 960212 | O3 | -81 | 16 to 9 | 36.9 | 2.9 |  | 0.08 | 0.008 | Strange profile above 32kM. Truncated above 32kM on .14E file to
|  | Z00841 | [8 empty] | redetermine total ozone. Box temp erratic with resistor. RH also increased
from 4% at 15kM to 33% at 24kM.
New ozone total 334.3 old value was 316.5 DU.
46 | 1996-02-13 | 960213 | O3 | -38 | 23 to | 7.2 | 398 |  | 0.07 | 0.016 | Failed at T+35 mins. Box temps badly in error for resistor flight. Journal file
|  | Z00842 |  |  | -3.3 | [5 empty] | remarks throughout flight - Error measuring sensor resistance
No LE. file as flight details not sent to NILU. Classic failure T,, U, BOX
TEMP then fail at 233mB.
47 | 1996-02-14 | 960214 | O3 | -79 | 14.5 to | 17.7 | 81 |  | 0.12 | 0.019 | Balloon burst early, too early for total ozone. Otherwise OK
|  | Z00843 |  |  | 16.5
48 | 1996-02-16 | 960216 | O3 | -100 | 14.5 | 32.9 | 5 |  | 0.1 | 0.013 | Numerous temp and RH's lost after 10mins but reappeared
|  | Z00844 | [8 empty] | temporarily at T+40 mins. Receiver jumped to Vaisala at T + 50 mins,
retuned.Temp looked OK from T+86 mins. Signal lost at T+118 mins
Journal file message.'Temp sample rejected' and also 'Thermistor probably
broken'. Signal low after 6 mins, see graph.MIN AIR TEMP OF -141 C
WRONG. Ozone OK and we got total ozone.
49 | 1996-02-18 | 960218 | O3 | -82 | 14.5 | 28.8 | 10.3 |  | 0.12 | 0.022 | Sonde battery lead broken off connector, resoldered. Otherwise OK
|  | Z00845 | [8 empty] | Our software detected launch time wrong. Otherwise OK
400' of descent data removed from .14e file. Change in total ozone 0.1DU
50 | 1996-02-19 | 960219 | O3 | -80 | 14.5 to | 31.4 | 6.8 |  | 0.05 | 0.01 | OK.
|  | Z00846 |  |  | 16.1
51 | 1996-02-21 | 960221 | O3 | -85 | 14.1 | 33.4 | 14.7 |  | 0.07 | 0.01 | OK but. disc missing, waiting for copy of zip file from Lerwick
Z00847
52 | 1996-02-23 | 960223 | O3 | -80 | 16.3 | 33.7 | 5.3 |  | 0.186 | 0.035 | OK. Our software has not detected launch properly
Z00848
53 | 1996-02-25 | 960225 | O3 | -77 | 13.8 to 19.0 | 27.3 | 14.1 |  | 0.07 | 0.02 | OK. No descent data however. Noise on box temp between 2kM and 3kM
|  | Z00849 |  |  | mode 14.5 | [5 empty] | correlates with TEMP, RH. Error computing sensor resistance JNL file
54 | 1996-02-26 | 960226 | O3 | -79 | 12.1 to 16.7 | 35.5 | 4 |  | 0.1 | 0.03 | OK. Big jumps in descent data and ozone data too low. Sudden drop in
|  | Z00850 | [8 empty] | ozone at top of ascent. Flight truncated at 34.7kM and ozone redetermined
new value 266.3 old value 259.5
55 | 1996-02-28 | 960228 | O3 | -77 | 13.8 to 17.1 | 26.5 | 17.5 |  | 0.12 | 0.02 | OK. RH 28% down to 13% in stratosphere. Some scatter in temp, RH
|  | Z00851 | [8 empty] | and box temp prior to launch. Wind was 30kts so it was getting knocked!
56 | 1996-03-01 | 960301 | O3 | -81 | 14.7 | 34.6 | 4.8 |  | 0.07 | 0.016 | OK. Some scatter in T, U and Box temp prior to launch. RH in stratosphere
|  | Z00852 | [8 empty] | 28% to 13%. 30kts at surface. RS-80 blotted out Viz for 1 min.
RS-80 400.92MHz _ VIZ 405.02MHz
57 | 1996-03-04 | 960304 | O3 | -85 | 14.5 | 33.8 | 5.1 | resistor | -0.02 | 0.037 | OK. Cell o/p wires reversed hence -ve currents. changed after monitoring
|  | Z00853 | [8 empty] | for launch
58 | 1996-03-05 | 960305 | O3 | -84 | 14.9 | 33.4 | 5.5 | resistor | 0.09 | 0.041 | OK.RH 99% at launch, 2% by 750 mB.
Z00854
59 | 1996-03-07 | 960307 | O3 | -76.4 | 14.7 | 34.2 | 5.5 | resistor | 0.09 | 0.025 | OK. No winds, LORAN aerial not unwound
Z00855
60 | 1996-03-08 | 960308 | O3 | -70.1 | 14.2 | 31.6 | 8.8 | resistor | 0.08 | 0.048 | OK. Fast descent. Improved parachute to be used.
Z00856
61 | 1996-03-11 | 960311 | O3 | -62 | 14.6 | 18.3 | 79.3 | resistor | 0.12 | 0.034 | Fail. Signal faded from 255 to 83 as frequency steadily rose from 404.9 to
|  | Z00857 | [8 empty] | 406.8 then jumped to 408.2 Mhz. Flight ended too early for full ozone
profile or total ozone
62 | 1996-03-13 | 960313 | O3 | -60 | 14.3 to | 33.5 | 6.9 | resistor | 0.13 | 0.02 | OK but box temp varies by 4 C
|  | Z00858 |  |  | 18.3
63 | 1996-03-15 | 960315 | O3 | -61 | 14.6 | 27.6 | 17.5 | resistor | 0.12 | 0.031 | OK. Frequency jumped from 404Mhz to 401Mhz at launch and several
|  | Z00859 | [8 empty] | times thereafter up to 5.5kM. OK after that.
64 | 1996-03-18 | 960318 | O3 | -58 | 23 to -6 | 34.2 | 8.1 |  | 0.21 | 0.038 | Missing data 5 mins after bead went -ve temps. T, U and bead reappeared
|  | Z00818 | [8 empty] | at intervals throughout flight.
Error computing sensor resistance on JNL file
0.16mA in ASC file.
65 | 1996-03-20 | 960320 | O3 | -61 | 14.7 to 16 | 30.9 | 9.4 | resistor | 0.11 | 0.023 | OK. Light winds and low signal strength (~140) because on omni aerial
|  | Z00861 | [8 empty] | still at T+60 mins. Great increase in signal when range increased and yagi
aerial in use
66 | 1996-03-25 | 960325 | O3 | -71 | 14.5 to
|  | z00862 |  |  | 15.5 | 32.5 | 7.7 | resistor | 0.23 | 0.047 | OK. Signal lost at burst
67 | 1996-03-26 | 960326 | O3 | -68 | 16 to
|  | Z00863 |  |  | 2.9 | 33.8 | 6.3 | resistor | 0.07 | 0.017 | FAILED. Ozone replaced by -99 or -ve values as box temp went to 2.9 C.
| 1996-03-29 | 960329 | O3 | -73 | 14.7 | 16.3 | 87 | resistor | 0.11 | 0.04 | FAILED. No data after 87mB. Last data on XS file at 138.2 mins, 8293 secs
Last data on 20e file at 5447 secs. Processed data missing intermittently
during flight
| 1996-03-30 | 960330 | O3 | -72 | 14.4 | 33.6 | 5.9 | resistor | 0.18 | 0.02 | Eventual success. Missing processed data throughout flight. Evidence of
RF interference. Background current increased from 0.017mA to 0.26mA
when ozonesonde installed in box. V poor signal strength on omni aerial
Repeated loss of signal due to system kept on jumping to aerial at 327
degrees when it should have been 120 degrees.
| 1996-03-31 | 960331 | O3 | -71 | 29 to 14 | 33.3 | 6.1 | resistor | 0.17 | 0.02 | Partial success. Box temp erratic and variable, with Error computing
sensor resistance in JNL file. Signal strength poor throughout. High
elevation and v. light winds. Lost of processed temp, RH and box temp at
about 220mB although RAW data still coming in. Automatic aerial selection
faulty, selecting 180 degrees instead of 120 degrees
| 1996-04-03 | 960403 | O3 | -72 | 14.6 | 32 | 7.8 | resistor | 0.16 | 0.015 | OK. System error 612. Program, antenna module time out on sending
request
| 1996-04-10 | 960410 | O3 |  |  |  |  | resistor |  |  | FAIL. Power supply failure. UPS did not hand back when mains reinstated
and eventually died, corrupting receiver software. Reinstalled under
instruction from VIZ
| 1996-04-17 | 960417 | O3 | -57 | 14.4 | 24.6 | 26.4 | resistor | 0.019 | 0.009 | OK. Active aerial and antenna mode missing from antenna display
Missing data mins 19 - 31 due to poor reception.
| 1996-04-24 | 960424 | O3 | -54 | 14.5 | 33.7 | 7 | resistor | 0.079 | 0.024 | OK. No Loran sync on french chain until after launch
| 1996-05-01 | 960501 | O3 | -50 | 16.5 | 15 | 117 | resistor | 0.035 | 0.077 | Failed. Continuous data from 2.1kM to 12kM then only a little more till
15kM. Signal strength poor dropping from 255 to 50 at fail. No winds on
WNX file until 3kM
| 1996-05-03 | 960503 | O3 | -51 | 16.4 | 10 |  | resistor | 0.098 | 0.024 | Fail at min 30 when pressure and ozone replaced by /////. RAW data and
pressure counts updating, signal strength and antenna direction OK until
signal fade at 107 mins. Only error in JNL file , sys error 612 severity 1.
Time out on sending request to antenna.
| 1996-05-08 | 960508 | O3 | -53 | 17.7 | 34.1 | 6.5 | resistor | 0.007 | 0.043 | OK
| 1996-05-17 | 960517 | O3 | -48 | 24.3 to | 16.1 | 101.6 | YES | 0.11 | 0.012 | FIRST WITH PTU MODIFIED WITH CAPACITOR TO REDUCE/
[5 empty] | 11.3 | [5 empty] | REMOVE RF INTERFERENCE. Failed early in flight when
Vaisala also flying. Light winds
| 1996-05-20 | 960520 | PTU | [7 empty] | Twin Flight failed. No radar. Initially it said receiver not matched, LORAN 1
LORAN 2 unmatched. Receiver downloaded. Cables OK
| 1996-05-21 | 960521 | PTU | [7 empty] | OK
| 1996-05-25 | [9 empty] | NEW RECEIVER AT LERWICK WITH CORRECTLY TUNED I. F.
PREVIOUS ONE WAS 1 MHz OUT AND CAUSING LOW SIGNAL
STRENGTHS BEING RECEIVED
NOTE max RF levels 216 approx hereafter c.w. 255 with old
receiver
| 1996-05-31 | 960531 | PTU | [7 empty] | 2 PTU flights OK
| 1996-06-05 | 960605 | PTU | [7 empty] | OK
| 1996-06-05 | 960605 | O3 | -53 | 16.3 | 31.8 | 9.9 | resistor | 0.135 | 0.02 | OK. JNL file says humidity sample rejected. Frequency out of range.
A few error 610 and 612 antenna errors.
0.13 mA chosen by system
| 1996-06-18 | [9 empty] | 9 modified Microsonde 2's (ozone) arrived Lerwick from Viz
| 1996-06-19 | 960619 | O3 | -51 | 24 to 12 | 33.3 | 8.1 | YES | 0.007 | 0.003 | OK. JNL file error 612 antenna time out on sending request. MODIFIED
| 1996-06-26 | 960626 | O3 | -53 | 26 to 14 | 34.4 | 7.1 | YES | 0.052 | 0.011 | OK. MODIFIED SONDE WITH CAPACITORS TO DRAIN AWAY RF
| 1996-07-03 | 960703 | O3 | -46.4 | 35 to 13 | 19.2 | 44.4 | YES | 0.17 | 0.023 | Failed at 19 kM. Background current measured and chosen by system
as 0.171 mA but measured by DVM as 0.023 mA amps.
JNL file errors 610,612,616. Signal strength at launch 219 ending at 126.
RAW data ends in IM2 file at T+9812 secs. WNX data ends at 4220 secs.
XS data ends at 9494 secs.
ASC file ozone out of range just prior to failure. MODIFIED SONDE
| 1996-07-10 | 960710 | PTU | -48 |  |  | 6.9 |  |  |  | OK. Curious jump of frequency from 400.5 MHz to 404.0 MHz after
1900 secs, associate with drop of signal strength from 230 to 130.
Switching to yagi aerial from omni increased signal strength from typically
135 to 161. This is as it should be.
| 1996-07-17 | 960717 | O3 | -60 | 28 to 20 | 34.3 | 7.1 | YES | 0.005 | 0.007 | OK. However at around launch and for approx 1 min into flight temp, box
[5 empty] | * | [5 empty] | temp out of range or too high. i.e. box temp at launch 163.1 C
air temp 29 C. Ground temp at launch 10.5 C. JNL file says Error
computing sensor resistance or temp sample rejected. This is indicative of
RF interference. MODIFIED SONDE
| 1996-07-31 | 960731 | O3 | -48 | 16.4 | 33.3 | 8.2 | NO | 0.12 | 0.025 | OK. Note difference between DVM current and system current.
SONDE NOT MODIFIED
| 1996-08-21 | 960821 | O3 | -44 | 19 to -23 | 11.7 | 207 | NO | 0.007 | 0.05 | FAIL. JNL file gives error computing sensor resistance. Box temp goes -ve
with resistor in place. T, %RH and bead temp processed data lost as bead
temp showed -5 C. RF interference problem. SONDE NOT MODIFIED
| 1996-08-23 | 960823 | O3 | -58 | 14.2 to | 34.2 | 4.7 | NO | 0.113 | 0.012 | OK. Note difference between DVM current and system current.
[5 empty] | 16.1 | [5 empty] | SONDE NOT MODIFIED
| 1996-08-28 | 960828 | O3 | -54 | 14.4 | 33.8 | 7.3 | NO | 0.065 | 0.044 | OK. Missing data on descent due to interference from Vaisala sonde
SONDE NOT MODIFIED
| 1996-09-11 | 960911 | O3 | -55 | 14.5 | 34.1 | 6.4 | NO | 0.009 | 0.009 | OK. SONDE MODIFIED BY D MOORE
| 1996-09-18 | 960918 | O3 | -61 | 13.5 | 31.3 | 9.6 | NO | 0.11 | 0.045 | OK. Priority 1 system errors 610, 612
OK. SONDE MODIFIED BY D MOORE
| 1996-09-27 | 960927 | O3 | -52 | 14.4 | 31.7 | 8.9 | NO | 0.009 | 0.062 | OK. New software at Lerwick, file names incorrect. Sorted for next flight
OK. SONDE MODIFIED BY D MOORE
| 1996-10-02 | 961002 | O3 | -44 | 27 to | 8.9 | 308 | YES | 0.056 | 0.007 | FAIL. At 300mB lost RAW and processed data.System errors 610,612,616.
[5 empty] | 17 | [5 empty] | rf levels 122 to 125 to 170. SONDE MODIFIED BY D MOORE
| 1996-10-09 | 961009 | O3 | -53 | 30 to 2 | 11.5 | 202 | YES | 0.008 | 0.07 | FAIL. Some -99 in ASC file. SONDE MODIFIED BY D MOORE
| 1996-10-16 | 961016 | O3 | -48 | 14.2 | 13.8 | 142 | NO | 0.018 | 0.0037 | FAIL on ascent, data available on descent. Sonde locate and match but
poor signal. Min 47 lost processed data. Min 78 lost RAW data. Correct
yagi used as per radar info. Sonde data ok again just before burst and on
descent. System kept on selecting aerial at 180 but radar showed 030.
4 lines of -99's for O3 and pressure, prior to losing data for 2nd time.
Around time of fail - Alarm condition 0 exists. Alarm turned on
SONDE MODIFIED BY D MOORE
| 1996-10-23 | 961023 | O3 | -67 | 12.9 to | 33.7 | 5.4 | NO | 0.05 | 0.012 | OK. Box temp noisy till launch. SYS errors 612,615,616, 850 (humidity
[5 empty] | 14.3 | [5 empty] | sample rejected) 0842 launch, Alarm turned on 11:12, 30 mins after end
of data
| 1996-10-30 | 961030 | O3 | -29 | 14.4 | 6.4 | 433 | NO | 0.04 | 0.014 | FAIL. JNL error 610,612,616. Alarm on at 09:55:19 fail 12 secs later
| 1996-11-06 | 961106 | O3 | -70.6 | 13.4 | 31.1 | 7.7 | NO | 0.018 |  | OK. Nursed along because of low signal strength. Red light appearing on
rear of Zeemet rack which went off when the unit was touched. Radar at
110 but aerial selected was 060. Techs found water in aerial switching unit
which was dried out. Pre-amp was changed for spare.Zeemet rack sent
from Viz. Spare plug in units sent from Herstmonceux. 4 ozonesonde
packages left. Approx 6 PTU ascents tried before 27/11/96 and a marked
improvement in signal strength found. Water in external boxes appears to
have been the cause.
| 1996-11-27 | 961127 | O3 | -79.2 | 13.4 | 33 | 5.5 | NO | 0.049 | 0.067 | OK. System errors 610 and 612. Marked improvement in signal strenth
PTU flights now working OK. Prior to pre-amp change and aerial switching
unit dried out they were also failing on poor signal near the Tropopause
| 1996-12-04 | 961204 | O3 | [7 empty] | Sudden failure at launch or just prior to it. 32kts mean wind.
AERIAL SWITCHING UNIT REMOVED AS IT WAS ACTING AS ATTENUATOR. IT WAS SEVERELY
CORRODED INSIDE WITH WATER GETTING IN PAST GASKET AND SOCKETS WITHOUT RUBBER GASKETS
MOST OF FLIGHTS NOW ON OMNI ONLY WITH STEERABLE YAGI SWITCHED IN MANUALLY LATE IN FLIGHT
SIGNAL STRENGTHS ABOUT 10 TIMES GREATER WITHOUT UNIT IN CIRCUIT.
| 1996-12-20 | 961220 | O3 | -74 | 13.6 | 32.8 | 5.8 | NO | 0.004 | 0.017 | OK. Numerous single line missing data ozone and height
| 1996-12-23 | 961223 | O3 | -72 | 13.4 | 33 | 5.7 | NO | 0.018 | 0.007 | OK.Loss of data 28.6kM to 31.8kM due to interference with Vaisala sonde
NEW BATCH OF SONDES FLOWN WITH THERMISTORS WITHOUT MODIFICATION
| 1996-12-25 | 961225 | O3 | -70 | 25 to - 5 | 33.2 | 5.9 | YES | 0.067 | 0..024 | OK.
| 1996-12-27 | 961227 | O3 | -64 | 24 to 9 | 22.2 | 36.6 | YES | 0.006 | 0.016 | OK. Early burst.
| 1996-12-30 | 961230 | O3 | -69 | 24 to 3 | 31 | 8.5 | YES | 0.029 | 0.024 | OK.
| 1997-01-01 | 970101 | O3 | -68 | 26 to 4 | 30.4 | 9.5 | YES | 0.006 | 0.016 | OK.
| 1997-01-03 | 970103 | O3 | -71 | 25 to 3 | 33.2 | 6 | YES | 0.005 | 0.024 | OK. Error computing sensor resistance @ 09:03:55. Code 6
| 1997-01-06 | 970106 | O3 | -73 | 26 to -12 | 42.7? | 1.2? | YES | 0.018 | 0.024 | OK. However pressure at 1 mB is out and so therefore is height at top
of ascent. Temp OK to 28kM but up to 6 C different from mid day ascent.
Box temp at -12 C unlikely but data looks OK
| 1997-01-08 | 970108 | O3 | -72 | 25 to 3 | 32.4 | 6.3 | YES | 0.012 | 0.005 | OK.
| 1997-01-10 | 970110 | O3 | -72 | 26 to -10 | 33.8 | 5 | YES | 0.016 | 0.003 | OK. However box temp down to -10 C on ascent -30 C on descent!!!!
| 1997-01-13 | 970113 | O3 | [7 empty] | Fail. No processed data
| 1997-01-15 | 970115 | O3 | -71 | 29 to 11 | 20.2 | 47.9 | YES | 0.017 | 0.005 | OK. Early burst
| 1997-01-17 | 970115 | O3 | -76 | 24 to 0 | 33.7 | 4.9 | YES | 0.032 | 0.015 | OK. 4 PTU's and 1 interface change before good combination. Humudities
missing near surface because values out of possible limits.
Same with 1 temp sample
| 1997-01-20 | 970120 | O3 |  |  |  |  | YES |  |  | Failed at Trop. No processed data. Launch never detected. Background
current drifting from 0.355 micro amps to 0.14 micro amps. PC very slow
throughout. Good signal strengths. Only 2 secs of flight data in ASC file
Alarm condition 0 exists, alarm turned on.
Plenty of RAW data. Flight could be reconstituted probably
| 1997-01-22 | 970122 | O3 | -60 | 28 to 8 | 12.1 | 184 | YES | 0.026 | 0.008 | Failed at trop with -ve ozone. Current suddenly went way off limits.
Alarm condition 0. Current suddenly way off limits as previous flight
| 1997-01-23 | 970123 | O3 | -60 | 31 to 7 | 12.2 | 176 | YES | 0.01 | 0.021 | Failed. Zero ozone at trop. IM2 file gives count of 2 instead of 30000 ozone
ozone current. Data at fail comes in pairs as part of data error check then
gap in time
Current suddenly way off limits as previous flight. System error 616
| 1997-01-24 | 970124 | O3 | -71 | 28 to 7 | 23.2 | 30.5 | YES | 0.031 | 0.002 | Ok, eventually. 6 Hours trying various combinations. We have 8 rejected
sondes now.
| 1997-01-27 | 970127 | O3 | -65 | 26 to 3 | 34.5 | 4.9 | YES | -0.028 | 0.028 | OK. No problems
| 1997-01-29 | 970129 | O3 | -61 | 26 to 5 | 30.9 | 9.5 | YES | 0.003 | 0.005 | OK. Comms errors and time out on sending request on Rx module
Numerous errors computing sensor resistance. Humidity sample
frequency out of range
| 1997-01-31 | 970131 | O3 | -63 | 29 to 5 | 14.7 | 123 | YES | 0.026 | 0.041 | Failed early in flight when ozone current RAW counts dropped to 2.
Processed current far too high causing failure. JNL file gives error 850
Humidity sample out of frequency range.
| 1997-02-02 | 970202 | O3 | -57 | 23 to 2 | 33.5 | 6.6 | YES | 0.034 | 0.071 | OK
| 1997-02-03 | 970203 | O3 | -62 | 25 to 1 | 32.9 | 6.8 | YES | 0.108 | 0.01 | OK
| 1997-02-05 | 970203 | O3 | -61 | 23 to -8 | 35.4 | 4.4 | YES | 0.111 | 0.023 | OK
| 1997-02-07 | 970207 | O3 |  |  |  |  | YES | 0.143 | 0.047 | Failed few secs after launch. JNL file sys error 616
JNL file errors. Response length does not match table. Alarm condition
0 exists, alarm turned on. CAP references rejected CHIGH 4767 <
CLOW 24670. Humidity sample rejected off limits. Launch detected
15:21:26 then alarm condition 0 exists. P, T, O3 data off limits just
after launch, appearing reasonable intermittently during flight
| 1997-02-10 | 970210 | O3 | -77 | 26 to 0.2 | 27.9 | 12.1 | YES | 0.082 | 0.048 | OK
| 1997-02-12 | 970212 | O3 | -73 | 13.5 | 26.8 | 15.5 | NO | 0.055 | 0.02 | OK. Resistor flight
| 1997-02-14 | 970214 | O3 | -60 | 25 to -2 | 28.7 | 12.2 | YES | 0.087 | 0.05 | OK. Sys error 616 around launch - length does not match table
| 1997-02-17 | 970217 | O3 | -68.8 | 25 to -2 | 33.7 | 5.4 | YES | 0.09 | 0.008 | OK. In JNL file 4 lines of alarm condition 0 exists, alarm turned on.
| 1997-02-18 | 970218 | O3 | -66 | 24 to 2 | 24.9 | 22.7 | YES | 0.05 | 0.009 | OK. Early burst
| 1997-02-19 | 970219 | O3 | -70 | 19 to -5 | 32.7 | 6.1 | YES | 0.044 | 0.009 | OK. Signal lost at 330kM range!
| 1997-02-20 | 970220 | O3 | -69 | 22 to -2 | 19.2 | 55 | YES | 0.061 | 0.028 | OK. Early burst JNL file sytem error few minutes prior to launch.
response length does not match table.
| 1997-02-21 | 970221 | O3 | -77 | 27 to -6 | 33.8 | 4.9 | YES | 0.102 | 0.082 | OK. JNL file system error 616, response length does not match table
| 1997-02-22 | 970222 | O3 | -79 | 24 to -7 | 26.6 | 15 | YES | 0.029 | 0.03 | OK. Box temp -12 on descent
| 1997-02-24 | 970224 | O3 | -75 | 22 to -22 | 25.9 | 16.1 | YES | 0.019 | 0.037 | OK. Box temp very cold and ozone lower on descent than ascent.
| 1997-02-25 | 970225 | O3 | -76 | 22 to -33 | 32.1 | 6.1 | YES | 0.045 | 0.051 | OK. Very cold box temp, extra insulation in box between battery and
pump. Can we believe these temps?
Flight alarm condition 0 exists, alarm turned on near end of flight.
| 1997-02-26 | 970226 | O3 | -67 | 27 to 4 | 25.6 | 18.9 | YES | -0.007 | 0.027 | OK. Many temp and RH rejected prior to launch. Also response length
does not match table. At end of flight alarm condition 0 exists alarm on
| 1997-02-28 | 970228 | O3 | -70 | 23 to -9 | 31.3 | 7.9 | YES | 0.007 | 0.026 | OK. Sys error 616 , time out on sending request once prior to launch
| 1997-03-03 | 970303 | O3 | -64 | 24 to -1 | 33.2 | 6.1 | YES | 0.06 | 0.023 | OK.
| 1997-03-05 | 970305 | O3 | -69 | 282 to | 21.7 | 38.9 | YES | 0.048 | 0.052 | OK. However box temp showing 200 C.
198
| 1997-03-06 | 970306 | O3 | -72 | 27 to -5 | 29.2 | 11.1 | YES | 0.008 | 0.023 | OK
| 1997-03-07 | 970307 | O3 | -76 | 24 to -1 | 32.2 | 6.6 | YES | 0.023 | 0.016 | OK. At end of flight - alarm condition 0 exists, alarm turned on
| 1997-03-08 | 970308 | O3 | -75 | 25 to -2 | 32.5 | 6.5 | YES | 0.003 | 0.012 | OK. Alarm condition 0 exists near end of flight in JNL file
| 1997-03-10 | 970310 | O3 | -67 | 25 to 6 | 34.7 | 4.8 | YES | 0.004 | 0.009 | OK.
No more comments about descent particularly near end of flight
| 1997-03-12 | 970312 | O3 | -63 | 27 to -5 | 35.6 | 4.4 | YES | 0.014 | 0.021 | OK.
| 1997-03-13 | 970313 | O3 | -66 | 21 to -1 | 37 | 3.5 | YES | 0.025 | 0.05 | OK. JNL file just prior to launch, humidity sample rejected also
some missing data
| 1997-03-14 | 970314 | O3 | -58 | 31 to 4 | 18.3 | 68.6 | YES | 0.023 | 0.033 | Failed when ozone when erratic then off limits at 18.3kM. Last hour of
flight alarm condition 0 exists, alarm turned on.
| 1997-03-15 | 970315 | O3 | -62 | 7 to -3 | 17.3 | 78.9 | YES |  |  | Failed with sudden loss of signal. ZIP and therefore JNL file unavailable
| 1997-03-17 | 970317 | O3 | -67 | 21 to -1 | 34.6 | 5.1 | YES | 0.017 | 0.037 | OK
| 1997-03-19 | 970319 | O3 | -63 | 24 to 6 | 34.7 | 5.1 | YES | 0.006 | 0.009 | OK.
| 1997-03-21 | 970321 | O3 | -61 | 32 to -9 | 34.5 | 5.4 | YES | 0.063 | 0.02 | OK. However box temp too low and JNL file has numerous temp sample
rejected unreasonable values
| 1997-03-24 | 970324 | O3 | -62 | 25 to 2 | 35.3 | 4.5 | YES | -0.011 | 0.022 | OK
| 1997-03-26 | 970326 | O3 | -63 | 24 to 3 | 30.8 | 9.2 | YES | 0.041 | 0.024 | OK
| 1997-03-28 | 970328 | O3 | -62 | 24 to 1 | 31.3 | 8.5 | YES | 0.012 | 0.023 | OK
| 1997-03-29 | 970329 | O3 | -69 | 23 to -13 | 28.9 | 11.5 | YES | -0.78 | 0.026 | OK. However background current way off, far too high c.w. DVM
| 1997-03-31 | 970331 | O3 | -73 | 23 to 1 | 32 | 7.2 | YES | 0.038 | 0.027 | OK. Signal temporarily lost during flight
| 1998-04-01 | 980401 | O3 | -64 | 29 to 0 | 32.4 | 7.3 | YES | -0.002 | 0.02 | OK.
| 1997-04-05 | 970502 | O3 | -66 | 23 to -8 | 30.9 | 9.2 | YES | 0.021 | 0.026 | OK
| 1997-04-08 | 970408 | O3 | -67 | 32 to -10 | 34.6 | 5.2 | YES | 0.023 | 0.028 | OK
| 1997-04-09 | 970409 | O3 | -68 | 26 to -1 | 30.9 | 8.7 | YES | 0.018 | 0.023 | OK. JNL file. 3 mins after arm for launch. Response length does not
match table
| 1997-04-11 | 970411 | O3 | -67 | 24 to 1 | 28.6 | 12.9 | YES | 0.025 | 0.036 | OK.
| 1997-04-16 | 970416 | O3 | -58 | 26 to 1 | 34.5 | 6 | YES | 0.028 | 0.015 | OK
| 1997-04-23 | 970423 | O3 | -61 | 26 to -5 | 35.7 | 4.7 | YES | -0.004 | 0.024 | OK
| 1997-04-30 | 970430 | O3 | -62 | 29 to 6 | 35.5 | 5.1 | YES | 0.002 | 0.035 | FAIL. Very noisy in stratosphere. NILU file retracted 7/5/97
| 1997-05-09 | 970509 | O3 | -57 | 27 to 6 | 32.9 | 7.7 | YES | -0.013 | 0.028 | Very poor signal. Lot of time adjusting yagi without success. Data missing
25.8km to 27.1km and 28.8 to 32.2. Interpolated data inserted and flight
truncated at 28km
| 1997-05-14 | 970514 | O3 | -53 | 28 to -2 | 32.9 | 8.2 | YES | -0.053 | 0.025 | radar burst 33095 metres
| 1997-05-25 | 970525 | O3 |  |  |  |  | YES |  |  | Experimental surface mounted technology version. Fail at launch
no ozone, pressure erratic
| 1997-05-26 | 970526 | O3 | -53 | 28 to 2 | 36.1 | 5.3 | YES | 0.128 | 0.023 | radar burst pressure 5.9mb. Experimental surface mounted technology
| 1997-06-11 | 970611 | O3 | -56 | 25 to 5 | 34.3 | 5.2 | YES | 0.008 | 0.027 | OK
| 1997-06-25 | 970625 | O3 | -51 | 27 to 4 | 33.9 | 7.5 | YES | 0.108 | -0.015 | OK
| 1997-07-09 | 970709 | O3 | -54 | 32 to 5 | 32.1 | 9.8 | YES | 0.008 | 0.037 | OK. Error computing sensor resistance between arm for launch and flight
| 1997-07-23 | 970723 | O3 | -58 | 27 TO -2 | 34 | 7.6 | YES | 0.013 | 0.024 | OK
| 1997-08-01 | 970801 | O3 | -54 | 28 to 10 | 34.6 | 9.7 | YES | 0.021 | 0.032 | OK. VAISALA. Not done in research mode for descent data
| 1997-08-06 | 970806 | O3 | -62 | 26 to 4 | 34.6 | 6.8 | YES | -0.07 | 0.024 | OK
| 1997-08-20 | 970820 | O3 | -61 | 25 to 3 | 25.1 | 26.6 | YES | 0.032 | 0.016 | OK. Early burst
| 1998-04-01 | 980104 | O3 | -63 | 25 to 2 | 32.3 | 7.4 | YES | 0.004 |  | OK. Some erratic box temps on ascent
| 1998-04-08 | 980408 | O3 | -57 | 23 to 4 | 30.9 | 9.9 | YES | 0.04 | 0.05 | OK
| 1998-04-15 | 980415 | O3 | -52 | 23 to -3 | 32.1 | 8.4 | YES | -0.05 | 0.027 | OK. -ve background current though.
| 1998-04-22 | 980422 | o3 | -56 | 31 to 14 | 32 | 8.7 | YES | -0.058 | 0.016 | OK.
| 1998-04-29 | 980429 | O3 | -51 | 26 to 1 | 33.3 | 7.4 | YES | 0.015 | 0.024 | OK. Last line of ascent deleted from 14E and LE files as glitch present.
| 1998-05-06 | 980506 | O3 | -52 | 27 to 5 | 32.1 | 8.7 | YES | -0.009 | 0.015 | OK. Software chosen max height 30975 metres. It was 32132 metres
| 1998-05-20 | 980520 | O3 | -58 | 26 to 0 | 29.9 | 12.5 | YES | 0.015 | 0.065 | OK.
| 1998-06-03 | 980603 | O3 | -50 | 24 to 2 | 32.4 | 8.9 | YES | 0.008 | 0.002 | OK. Large spike on start of descent.
| 1998-06-17 | 980617 | O3 | -48 | 31 to 17 | 23.5 | 9.2 | YES | 0.015 | 0.005 | OK. Final box temp unusually high. Db 10.2 C
| 1998-07-01 | 980701 | O3 | -50 | 27 to 5 | 31.2 | 11.1 | YES | -0.027 | 0.02 | OK.
| 1998-07-15 | 980715 | O3 | -49 | 27 to 9 | 31.7 | 10.4 | YES | 0.005 | 0.018 | OK.
| 1998-07-29 | 980729 | O3 | -47 | 27 to 9 | 34.4 | 7.1 | YES | 0.018 | 0.005 | OK.
| 1998-08-12 | 980812 | O3 | -50 | 28 to 11 | 34.6 | 6.7 | YES | -0.03 | 0.026 | OK.
| 1998-08-26 | 980826 | O3 | -50 | 21 to -4 | 34.4 | 6.8 | YES | -0.046 | 0.022 | OK. Bad ozone data 25.4kM to 26.9kM. Repaled by interpolated data
prior to sending to NILU. Ozone total using amended file. 14E and 20E
files also amended.
| 1998-09-09 | 980909 | O3 | -55 | 26 to 10 | 33.3 | 7.5 | YES | -0.019 | 0.018 | OK. Some rejected out of range data just prior to launch.
-ve background current.
| 1998-09-23 | 980909 | O3 | -61 | 27 to 10 | 34.6 | 5.9 | YES | 0.005 | 0.02 | OK
| 1998-10-09 | 981009 | O3 | -61 | 27 to 10 | 33.9 | 7.5 | YES | 0.005 | 0.02 | OK
| 1998-10-21 | 981021 | O3 | -64 | 28 to 9 | 35.9 | 4.1 | YES | 0.022 | 0.018 | OK
| 1999-04-28 | 990428 | O3 | -64 | 28 to 0 | 30.7 | 10.3 | YES | -0.016 | 0.018 | OK. Amended near burst as sudden jump in ozone
| 1999-05-05 | 990505 | O3 | -60 | 29 to 15 | 15.6 | 109 | YES |  | 0.018 | Fail. Error computing sensor resistance. RF problems
| 1999-05-07 | 990507 | O3 | -63 | 31 to 17 | 27.9 | 16 | YES | 0.135 | 0.012 | OK. However evidence of RF interference in background curreent.
Also 14e file edited to remove last 2 lines. Errors in ozone, sudden jump
LE file OK anyway
| 1999-05-19 | 990519 | O3 | -59 | 26 to 0 | 23.6 | 32.3 | YES | 0.026 |  | OK, but sudden failure at 23.6kM. PC locked up
| 1999-06-02 | 990602 | O3 | -55 | 29 to 9 | 31 | 10.8 | YES | 0.008 | 0.015 | OK
| 1999-06-16 | 990616 | O3 | -55 | 28 to 3 | 31 | 11.1 | YES | -0.032 | 0.039 | OK
| 1999-06-30 | 990630 | O3 | -49 | 30 to 11 | 31 | 11.4 | YES | 0.024 | 0.059 | OK
| 1999-07-14 | 990714 | O3 | -48 | 26 to 6 | 34.9 | 6.6 | YES | -0.05 | 0.029 | OK
| 1999-07-28 | 990728 | O3 | -62 | 26 to 6 | 33.3 | 8.3 | YES | 0.022 | 0.05 | OK.
NO MORE OPERATIONAL SONDES TO FLY

## Summary
SUMMARY OF VIZ OZONESONDE FAILURES SINCE FIRST SONDE FLOWN 6/11/95 till Apr 97
PERCENT
DATE | SUCCESS | FAIL | SUCCESS | MAIN CAUSE OF FAILURES OR PROBLEMS DURING FLIGHTS
1995-11-01 | 8 | 3 | 72.7273 | Missing processed data, unable to process bead temp, temp and RH out of range
1995-12-01 | 7 | 2 | 77.7778 | Missing processed data, temp, box temp and RH out of range, error computing sensor resistance
1996-01-01 | 12 | 2 | 85.7143 | Red light on rear panel or total loss of signal. Missing data on other "successful" flights
1996-02-01 | 16 | 1 | 94.1176 | Unable to process box temp, temp, RH because out of range. RAW data continues to come in unprocessable
1996-03-01 | 10 | 3 | 76.9231 | Missing processed data, temp, box temp and RH out of range, error computing sensor resistance
1996-04-01 | 3 | 1 | 75 | Poor signal and data out of range so unable to be processed.
1996-05-01 | 1 | 3 | 25 | Poor signal, data out of range. Pressure and ozone replaced by ////. No winds in WNX file.
1996-06-01 | 3 | 0 | 100 | Some humidity data out of permissible range. A few 610, 612 antenna error messages
1996-07-01 | 3 | 1 | 75 | Poor signal and frequency shift. Box temps in error. Background current inaccurate
1996-08-01 | 2 | 1 | 66.6667 | Box temp -ve even with fixed resistor substitute. System couldn't process bead temp. temp and RH
1996-09-01 | 3 | 0 | 100 | OK. A few 610, 612 errors in JNL file
1996-10-01 | 1 | 4 | 20 | Aerial problems, incorrect aerial selected. Poor signal. Data out of range so couldn't process JNL, 610,612,616
1996-11-01 | 2 | 0 | 100 | Flights nursed along, poor signal. Water in aerial switching unit.
1996-12-01 | 5 | 1 | 83.3333 | Sudden failure at launch
1997-01-01 | 10 | 5 | 66.6667 | Box temps too low, failure at trop -ve ozone, sudden counts of 2 in IM2 file for ozone. Launch undetected
1997-02-01 | 16 | 1 | 94.1176 | Box temps too cold down to -22 C. Error 616 in JNL file
1997-03-01 | 16 | 2 | 88.8889 | Ozone too low, box temp 200 C!!
1997-04-01 | 3 | 0 | 100
TOTAL | 121 | 30 | 80.1325 | Problems were still being experienced even on successful flights.
In 1997 to
1997-04-08 | 45 | 8 | 84.9057
Ht starts at 122M. OK Bead thermistor replaced by resistor. Temp
unstable after burst
Missing data between 190 and 25mB. Temp, Rh Box temp replaced by //
not transmitted. 2nd sonde OK. Interference from Viz on deck to Vaisala
in flight